사 양 요 구 서

부서명 : 신경외과(통증클리닉)

물품분류 번 호	품목번호	품명 (Descriptions)	단위	수량
		초음파기 (Ultrasound system)	set	1

1. 일반적인 요구사양

1. 무상써비스기간이 3년 이상인 제품(트랜듀서포함)

2. 본원의 PACS와 연동이 가능한 제품

3. 220V, 60Hz의 전원을 사용하는 제품

2. 사 양

1. Main Console

1) System Operating/Display Modes

- 초음파 발생 방식 : Digital Beamformer

- B-Mode, M-Mode가 가능한 제품

- Color Flow mode가 가능한 제품

- PW Doppler with High PRF(PW)가 가능한 제품

- Color Flow Imaging(CFI) 혹은 Color Doppler images가 가능한 제품

- Power Doppler Imaging(PD) 혹은 PDI가 가능한 제품

- Simultaneous B/PW, B/CFM, B/PD, B/TD, B/M가 가능한 제품

- Colorized B, M, and PW Doppler display가 가능한 제품

- Imaging frequency range : 1.5 - 18 MHz 혹은 2 ~16 MHz인 제품

2)  B-Mode and M-Mode Imaging

- Gray scale levels : application에 따라 자동조절 혹은 256 level인 제품

- Persistence : 7 steps이상인 제품

- B Gray Scale Map : application에 따라 자동조절 혹은 13 maps인 제품

- User selectable scan depth : 1 cm - 27 cm 혹은   0cm - 30 cm인 제품

- B Colorization : application에 따라 자동조절 혹은 16 chroma maps인 제품

3) Color Flow Mode 혹은 Color Flow Imaging

- Frequency Range : application에 따라 자동조절 혹은 Pen, Gen(depending

on probe)

- Color Maps : application에 따라 자동조절 혹은 8 Maps인 제품

- CFM Gain : 3 steps 혹은 0 ~ 100 %인 제품

- CFM Velocity Scale Range

PRF : 3 steps, probe dependent 혹은 600 Hz to 14k Hz인 제품

- CFM Spectrum Inversion : Invert On / Off

- CFM Baseline Shift : application에 따라 자동조절 혹은 8 steps인 제품

4) Power Doppler Imaging(PD)

- Frequency Range : application에 따라 자동조절 혹은 Pen, Gen(depending

on probe)인 제품

- PD Maps : application에 따라 자동조절 혹은 8 Maps인 제품

- Gain : 3 steps 혹은 0~100%인 제품

- Velocity Scale Range

PRF : 3 steps, probe dependent 혹은 1 ~ 23khz인 제품

5) Measurements and Calculations

- General B-Mode :  Distance(Point to Point), 2D Trace(Trace to length)

- Area/Circumference : Ellipse/ Trace

6)  Image /Volume Storage

- Cine memory : 120 MB Cine memory 혹은 Image cine : 5400 frames

loop cine : 8192 lines

- Image file : DICOM file, JPEG file, MPEG4 혹은 DICOM file, JPEG file

인 제품

7) Operator Interface

- Active probe ports : 1 port 이상인 제품

- 조작방법 : 터치스크린 혹은 키보드인 제품

- 모니터 크기 : 10.4인치 이상인 제품

- 해상도 :  800 x 600이상인 제품

2. 이미지 최적화 기능

1) 원터치이미지 최적화기능 : Auto Optimization 혹은 QuickScan

2) Coded Harmonics 혹은 Tissue Harmonic

3) Raw Data Processing : Post Processing 혹은 Post Imaging Processing

4) Speckle Reduction Imaging : SRI 혹은 SRF

5) Frequency Focus Composite : ATO (Auto Tissue Optimization) 혹은

Dynamic Receive Focusing

6) CRI( CrossXbeam ) 혹은 SCI

7) HD-Flow( High Definition Flow) 혹은 DPD ( Directional Doppler )

3. 트랜듀서

1) Wide Band Convex Transducer

- Frequency : 2.0 - 5.5 MHz 혹은 2.0 - 6.0 MHz

- Application :  Abdomen, Musculoskeletal, Thoracic / Pleural 혹은

Abdomen, OB Gyn, Urology, Pediatrics. Musculoskeletal

2) Wide Band Linear Transducer

- Frequency :  5.0 - 13.0 MHz인 제품

- Application : Vacular Access, Anesthesia, Point of care, Interventional

and MSK 혹은 Breast, Small Parts, Vascular, Musculoskeletal, Pediatrics

4. DICOM3.0 Connectivity

- 본원의 PACS와 연동이 가능한 제품

- DICOM Store / Worklist Query가 가능한 제품

3. 구성품

A) 혹은  B)

A) 구성품

1. 초음파진단기 본체 ------------------------------------1 Unit

- Tablet Style, Lithium-Ion Battery Pack with 1 probe port, Speaker

- Anesthesia, Intervention, Musculoskeletal, Point of Care and

Vascular Access Application

- Operating Modes - B-Mode, M-Mode, Color Flow Mode(CFM),

Power Doppler Imaging (PDI)

- Touch Screen

- Touch and Stylus user interface

- 10.4 inch High-Resolution Color LCD-Display: 800x600

- 170 degree wide angle view

- Standard CINE Memory (120MB)

- Removable secure digital storage (SD)

- Loops storage from memory

- Automatic Optimization - Auto Tissue Optimization :ATO

- Distance Measurement

- Pre-Processing: Acoustic Power Output, Color Flow Mode

- Post-Processing: Auto Tissue Optimize (ATO)

- Software Intensive Ultrasound Imaging Platform

∙ Digital Beamformer

∙ Continuous Dynamic Receive Focus/Aperture

∙ Multi-Frequency/Wideband Technology

∙ Image Reverse: Rihgt/Left

- Image Archive/Connectivity

- DVI Interface in docking station and docking cart

- USB Interface on docking station and docking cart

- Docking Connector

- Operating Manual

2. Probe Options

- 12L-SC Wide Band 4 - 13 MHz Linear Probe ----------------- 1 Set

- 4C-SC Wide Band 1.8 - 5.0 MHz Convex Probe --------------- 1 Set

3. USB thermal B&W printer: Sony UPD-897 ------------------- 1 Set

4. Printer Installation Kit ------------------------------------1 Set

5. USB Wireless Adapter ------------------------------------- 1set

6. Docking Cart BT11 --------------------------------------- 1set

7. DICOM Connectivity ------------------------------------- 1 Set

B) 구성품

1. MAIN ( IMAGING UNIT INCLUDED) SYSTEM -------- 1set

- CPU MODULE PENYIUM IV 2GHz  PROCESSOR

- 4 ACTIVE ARRAY PORST

- REAL TIME 129,024 CHANNEL SYSTEM

- MULTI BEAM RECEIVING SYSTEM

- Synthetic Aperture Control

- Motion mode (M-mode)

- Color Motion mode (CM-mode)

- Color Doppler Imaging

- Power Doppler Imaging

- Directional Power Doppler

- Pulsed wave (PW) spectral Doppler

- Multi frequency Doppler

- Tissue Harmonic Imaging

- Pulse Inversion Harmonic Imaging

- Tissue Doppler Imaging

- Extreme High Dynamic Range (170dB)

- Full Spectrum Imaging

- Power Pulse Inversion Harmonic Imaging

- Trapezoidal Imaging

- Color Edge Processing

- Combined modes

- Cine for 5,400 frames and Loop Review for 8,192 lines

- Clip Cine Store

- Volume Cine Save(128MB)

- Color VGA monitor with non-interlaced display

- Orientation CubeTM

- Integrated and DVD RW(CD R/W) drive & USB2.0(10 ports)

drive

- 2D QuickScanTM

- Doppler QuickScanTM

- Real-Time Auto Calculation

- Free Angle M-Mode

- 17“LCD VGA Color Monitor (상하, 좌우360도 전환)

- SCI(Spatial Compoud Image)

2. DICOM 3.0 ------------------------------------- 1EA

3. PROBE

1) C2-5EL PROBE (Convex) -------------------------- 1ea

2) L5-13IS PROBE (Linear) -------------------------- 1ea

4. BW Printer --------------------------------------1EA

5. OPERATION & Services MANUAL ----------------- 각1EA

6. DynamicMR plus  -------------------------------- 1EA

4. 기타사항

1. 상기 모든 내용을 모두 충족시키는 제품을 납품하여야 하며 납품 후에라도 상기

요구 사항에 맞지 않는 제품을 납품하였을 때에는 장비의 회수 및 결제 대금을 전액

변상하여야 하며 병원에 입은 손해에 대하여 보상하여야 한다.

2. 장비의 검수는 당 의료원의 의공담당자의 참석하에 실시 한 것만을 인정한다.

3. 장비설치 및 납품시 장비의 고장 및 불량이 발생하였을 경우 즉시 신장비로 교체   하여 재 납품 하여야 한다.

4. 무상 및 유상 써비스 기간내에 원외반출 수리시에는 대체장비를 지원하여야 한다.

5. 장비납품 및 설치시 해당병원 장비별로 납품․설치하여야 하며, 정상운영 될 때까

지 초기교육을 실시하며 교육과 관련한 비용 및 장비설치에 따른 제반비용을 부담하여

야 한다.

6. 장비의 사용자의 변경이 생겼을 경우 Operator 교육을 실시하며 그 제반비용은 무상으

로 처리한다.

7. 무상 A/S기간 중 장비운영과 관련한 모든 프로그램의 Up-grade시 장비 제조사로부터

무상 제공하여야 한다.

8. 제품의 포장은 검수시에 개봉한 것만을 인정한다.

9. 납품은 병원이 요구하는 시간과 장소에 납품한다.

10. 장비의 검수는 장비설치 및 셋팅 완료 후에 실시한다.

11. 장비의 무상보증 기간은 장비 설치 및 검수 후부터 시작한다.

12. 장비 설치시 병원의 설비의 변동이 없이 설치하는 것을 원칙으로 하되 부득이하게

공사 혹은 구조변경이 필요한 경우 병원에 사전에 통보하여야 하며 공사 시공시 병원

의 요청대로 시공하여야 한다. 또한 공사비용은 장비의 납품업체에서 부담한다.