사 양 요 구 서

부서명 : 영상의학과

물품분류 번 호	품목번호	품명 (Descriptions)	단위	수량
		초음파기 (Ultrasound system)	set	1

1. 일반적인 요구사양

1. 영상의학과에서 사용하기에 최적화되어 있는 제품

2. 트랜듀서 연결포트가 3개 이상인 제품

3. 무상써비스기간이 3년 이상인 제품(트랜듀서포함)

4. 본원의 PACS와 연동이 가능한 제품

5. 모니터의 position 변동이 쉬운 제품

6. 220V, 60Hz의 전원을 사용하는 제품

7. 사용자 편리하도록 설정이 가능한 제품(원터치스위치 혹은 터치 패널)

8. 모니터 크기는 19인치 이상인 제품

9. 원터치로 이미지의 최적화가 가능한 제품

10. CD/DVD로 데이터 백업이 가능한 제품

2. 사 양

1. Applications

1) Abdominal

2) Small Parts

3) Breast

4) Vascular 등

2. Imaging Modes

1) 2D grayscale imaging

2) Adaptive Broadband flow imaging 혹은 Color Doppler Velocity &

Eneragy 혹은 Color Flow Imaging

3) SonoCT Real-time Compound Imaging 혹은 SieClear spatial

Compounding 혹은 CrossBeam Imaging (X Beam)

4) XRES Adaptive Image Processing 혹은 Dynamic TCE 혹은 SRI( Speckle

Reduction Imaging)

5) Contrast Specific Imaging (CSI) 혹은 CPS with ADI 혹은 Coded Contrast

Imaging

6) Tissue Harmonic Imaging

7) Tissue Harmonic SonoCT Imaging 혹은 Extend Imaging Technology 혹

은 Coded Harmonic Imaging

8) Color Power Angio Imaging(CPA) 혹은 CDE and CDV 혹은 Power

Doppler Imaging

9) ISCAN with AGC 혹은 TEQ(+Auto doppler optimize) 혹은 Auto

Optimization (AO)

10) Pulsed Wave Doppler

11) High Definition Zoom mode

12) Reconstruction zoom with pan

13) Dupplex and simultaneous 2D/ PW Doppler

14) Independent  Triplex for simultaneous 2D/Color/PW

15) Vascular Enhancement 혹은 Clarify VE Techonolgy 혹은 Voxel Image

Processing Package

16) High Q control 혹은 Auto Doppler 혹은 Accumulation Mode

17) Panoramic Image 혹은 SieScape panoramic Imaging(color가능) 혹은

LOGIQ View

18) Advanced Breast Package 혹은 Fatty Tissue Imaging 혹은 Raw Data

analysis

19) Elastography 혹은 eSie Touch elasticity imaging 혹은 Elastography

20) Elasto-Quantification 혹은 S2000 eSie Strain Ratio

21) Elasto-point quantification (Elasto sharewave PQ) 혹은 Virtual Touch

Tissue Imaging 혹은 동등한 기능

3. Gray Shades

- 256(8 bits) in 2D, Doppler spectral analysis

- 64(6 bits) in M-mode

4. Acquisition frame Rate : 500f/s or more

5. System Architecture

1) dynamic range : 180dB 혹은 210dB 혹은 145dB

2) digital channels : 57,000channels 혹은 67,392 channels 혹은 56,000

channels

3) Multi-application workflow Protocols

6. Image Presentation

1) Up / down

2) Left / right

3) Multiple duplex image formats (1/3-2/3, 1/ 2-1/2,  2/3-1/3 ,50/50 and

full screen )

7. Adaptive Broadband flow imaging 혹은 Color Doppler Velocity & Eneragy

혹은 Color Flow Imaging

8. Contrast Specific Imaging (CSI) 혹은 CPS with CTI 혹은 Coded Contrast

Imaging

9. Tissue Harmonic Imaging

10. Tissue Harmonic SonoCT Imaging 혹은 Extend Imaging Technology 혹은

Coded Harmonic Imaging

11. Color Power Angio Imaging(CPA) 혹은 CDE and CDV 혹은 Power

Doppler Imaging

12. Measurement Tools and General Description

- 2D distance

- 2D circumference/ area by ellipse, continuous trace, trace by points

- 2D curved-linear distance

- Manual Doppler distance and trace

- Doppler values including PI,RI,S/D indices

13. Cineloop Review

- Acquisition, storage in memory and display in real-time and duplex

Modes of up to 2,200 frames or more

14. ISCAN with AGC 혹은 TEQ (Doppler포함) 혹은 Auto Optimization (AO)

- One - touch image optimization 2D and Doppler

- Operation in conjunction with SonoCT and XRES imaging Adaptive Gain

Compensation (AGC)

15. IFOCUS Intelligent Focusing 혹은 Dynamic Transmit Focusing 혹은

Ultra-Coded Excitation

16. IOPTIMIZE Intelligent Optimization 혹은 DIMAQ imaging optimization

혹은 Voxel Image Processing Package

17. SonoCT Real-time Compound Imaging 혹은 SieClear spatial

Compounding 혹은 CrossBeam Imaging (X Beam)

18. XRES Adaptive Image Processing 혹은 Dynamic TCE 혹은 SRI-HD

19. Advanced Breast Package 혹은  Fatty tissue Imaging 혹은 Elasto -

Breast

20. Tissue Aberration Correction 혹은  Fatty tissue Imaging 혹은 Virtual

Convex Imaging

21. Elastography 혹은 eSie Touch elasticity imaging

22. Elasto-Quantification 혹은 S2000 eSie Strain Ratio

23. Elasto-point quantification (Elasto sharewave PQ) 혹은 Virtual Touch

Tissue Imaging 혹은 동등한 기능

24. DATA Intelligent Management 혹은 S2000 Integrated Ultrasound

Workstation 혹은 Raw Data Anaalysis(post-processing)

25. Auto IMT 혹은 AHP

26. 단면적 자동 측정기능 : Measurement tools and general description 혹은

eSie Cal 혹은 Measurements & Calculations package

27. Connectivity  or DICOM 3.0 Storage /Print / Worklist

1) Standard connectivity features

- DICOM Print

- Image and Waveform export to removable media(DVD/CD)

- Export of report data to off-line analysis computer programs

- 100 base-T Ethernet output

2) NetLink Connectivity Option

- Image and waveform export to network storage servers

- DICOM Worklist with RIS support and automatic patient demographic

entry

- Performed Procedure Step(PPS)

- Storage Commit(SC)

- Structured Reporting(SR) includes OB/GYN and vascular

3) DiskLInk-DVD

- The system supports standards DVD and CD media, and allows

the following capabilities:

• Read and Write(once) to CD(CD+R)

• DVD Read + Write(DVD + RW)

- Data storage formats include DICOM, JPEG and AVI for motion clips

- User selectable compression is available for motion clip storage

- DICOM images stored on disc can be recalled on the ultrasound system

- JPEG images and AVI clips can be recalled on off-line viewing stations

4) On-board Patient Exam Storage

- Direct digital storage of B/W and color loops to internal hard Disk

drives

- Combined 160 GB storage capacity or more

3. 구성품

A) 혹은 B) 혹은 C)

A)구성품

1. Basic architecture ---------------------------------------- 1sys

- xSTREAM system architecture

- Broadband acoustic beamforming

- 57,000 digital channels

- 180 dB dynamic range

- 1TB hard disk drive

- integrated CD/DVD

2. Imaging Modes ------------------------------------------- 1set

-2D grayscale imaging with advanced pulse coding, pulse sharping,

-Multivariate Tissue Harmonic Imaging

-Tissue aberration correction

-Coded beamforming with chirp transmit

-Next Generation SonoCT beam-steered real-time compounding imaing

-Harmonic sonoCT imaging

-Next Generation XRES adaptive image processing technology

-ISCAN intelligent scanning with one- button TGC, gain and compression

map optimization

-Broadband flow imaging

-Philips Color Power Angio Imaging(CPA)

-Color compare mode

-Freehand 3D and MPR imaging

-Contrst imaging

3. Adaptive broadband Flow imaging --------------------------- 1ea

4. Contrast Imaging  ---------------------------------------- 1ea

5. Tissue Harmonic Imaging ---------------------------------- 1ea

6. Color Power Angio Imaging (CPA) --------------------------- 1ea

7. Freehand 3D volume and MPR imaging ----------------------- 1ea

8. ISCAN Intelligent Optimization ------------------------------ 1ea

9. IFOCUS Intelligent Focusing technology ----------------------- 1ea

10. IOPTIMIZE Intelligent Optimization ------------------------- 1ea

11. SonoCT Real-time Compounding Imaging --------------------- 1ea

12. X-RES Imaging ----------------------------------------- 1ea

13. Advanced Breast Imaging --------------------------------- 1ea

14. Tissue Abberation Correction ------------------------------ 1ea

15. Panoramic Imaging -------------------------------------- 1ea

16. Elastography ------------------------------------------- 1ea

17. Elasto-Quantification(EQ) --------------------------------- 1ea

18. Elasto-point quantification (Elasto sharewave PQ)--------------- 1ea

19. Auto IMT ---------------------------------------------- 1ea

20. Measurement tools and general description -------------------- 1ea

21. DATA Intelligent Management ----------------------------- 1ea

22. DICOM Connectivity ------------------------------------- 1ea

23. Clinical S/W ------------------------------------------- 1set

24. C5-1 Broadband PureWave Crystal Transducer --------------- 1ea

- Abdominal applications

25. L12-5 Broadband Linear Array Transducer ------------------- 1ea

- Small Parts & Superficial & MSK application

26. L17-5 Broadband Linear Array Transducer ------------------- 1ea

-  MSK & Superficial application

27. C9-5ec Broadband Curved Array Transducer ------------------ 1ea

-  Endo rectal application

28. Operation & Service Manual ------------------------------- 1set

29. BW printer 및 color printer ----------------------------- 각 1ea

30. 기타 장비구동에 필요한 악세사리 ----------------------------- 1set

B) 구성품

1. 본체 --------------------------------------------------- 1ea

2. 어디벤스트 지클리어 공간합성기능  ----------------------------- 1ea

3. 2D 그레이 스케일 이미징 ------------------------------------ 1ea

4. 칼라도플러 ----------------------------------------------- 1ea

5. 하나피 렌즈 기술 ------------------------------------------ 1ea

6. 내장형 S2000 초음파 워크스테이션 ----------------------------- 1ea

7. S2000 패키지 -------------------------------------------- 1ea

8. 다이나믹 트렌시트 포커싱  영상모드 ---------------------------- 1ea

9. eSie Touch Elasticity Imaging(엘라스토이미지) ------------------ 1ea

10. eSie Strain Ratio(엘라스토  래이시오 표시) --------------------- 1ea

11. 다이나믹 티슈 콘트라스트 인헨스먼트 -------------------------- 1ea

12. 티슈 이퀄리제이션, TEQ ----------------------------------- 1ea

13. 이지 칼큘레이션(오토area측정) ------------------------------- 1ea

14. 어드밴스 브레스트 이미징 ----------------------------------- 1ea

15. Clarify VE(혈관 다중허상 제거) ------------------------------ 1ea

16. AHP(Auto IMT)------------------------------------------ 1ea

17. Virtual Touch Tissue Imaging ------------------------------ 1ea

18. 19" 고 해상도 LCD 모니터 --------------------------------- 1ea

19. 6C1 HD  ----------------------------------------------- 1ea

20. 18L6HD ----------------------------------------------- 1ea

21. 9L4 --------------------------------------------------- 1ea

22. EC9-4 ------------------------------------------------ 1ea

23. Operation & Service Manual ------------------------------- 1ea

24. BW printer --------------------------------------------- 1ea

25. DICOM Connectivity ------------------------------------- 1ea

26. BW printer 및 color printer ----------------------------- 각 1ea

27. 기타 장비구동에 필요한 악세사리 ----------------------------- 1set

C) 구성품

1 . System architecture -------------------------------------- 1set

- Agile Acoustic Architecture

- 19" High Resolution LCD monitor

- 10.4" High resolution, Color, Touch LCD Screen

- 4 Active Probe Ports

- Floating keyboard adjustable in three(Height & Rotation & Extension)

demensions

- Power-assisted keyboard up-down control

- Integrated 160GB hard drive

- Integrated DVD/CD-RW drive

- Integrated multiple USB port

- Integrated speakers with sub-woofer for premium sound

- Integrated Gel warmer

- 256MB CINE Momory

- Export JPEG, WMV(MPEG4), and AVI form

2. Imaging Mode ------------------------------------------- 1set

- B-mode

- M-mode

- Color Flow Mode

- Power Doppler Imaging(PDI)

- PW Doppler

- Volume Modes (3D)

- Anatomical M-mode

- Coded Harmonic Imaging

- Coded Contrast Imaging

3 . CrossXBeam (Spatial Compounding) ------------------------- 1ea

4. SRI-HD (Speckle Reduction Imaging) ------------------------- 1ea

5. Raw Data Analysis --------------------------------------- 1ea

6. Coded Harmonic Imaging (CHI) ----------------------------- 1ea

7. Advanced 3D -------------------------------------------- 1ea

8. Automatic Optimization (AO)  ------------------------------- 1ea

9. Anatomical M - Mode ------------------------------------- 1ea

10. Accumulation Mode -------------------------------------- 1ea

11. Measurements & Calculations package ----------------------- 1ea

- Real-time Doppler Auto Measurements / Calculations

- OB Measurements / Calculations

- GYN Measurements / Caluculation

- Vascular Measurements / Calculations

- Urological Calcs

12. Virtual Convex ------------------------------------------ 1ea

13. Fine Angle Steer ---------------------------------------- 1ea

14. Built IN Gel Warmer ------------------------------------- 1ea

15. Matrix Transducer(Voxel Image Processing) Technology --------- 1ea

16. DICOM3.0 Connectivity  ---------------------------------- 1ea

17. Auto IMT ---------------------------------------------- 1ea

18. Elastography & Elasto-Quantification ----------------------- 1set

19. ML6-15-D Matrix Array Linear Probe ----------------------- 1ea

1) Application : Small parts, Vascular, Neonatal, Pediatrics

20. 9L-D Linear Probe -------------------------------------- 1ea

1) Application : Vascular, Small parts, Pediatric, Abdomen

21. C1-5-D Convex Probe ------------------------------------ 1ea

1) Application : Abdomen, OB Gyn, Urology, Vascular

22. IC5-9-D Micro-Convex Probe ------------------------------ 1ea

1) Application :OB Gyn, Urology

23. BW printer 및 color printer ------------------------------ 각1ea

24. Operating & Service Manual ------------------------------ 1set

25. 기타 장비구동에 필요한 악세사리 ---------------------------- 1set

4. 기타사항

1. 상기 모든 내용을 모두 충족시키는 제품을 납품하여야 하며 납품 후에라도 상기

요구 사항에 맞지 않는 제품을 납품하였을 때에는 장비의 회수 및 결제 대금을 전액

변상하여야 하며 병원에 입은 손해에 대하여 보상하여야 한다.

2. 장비의 검수는 당 의료원의 의공담당자의 참석하에 실시 한 것만을 인정한다.

3. 장비설치 및 납품시 장비의 고장 및 불량이 발생하였을 경우 즉시 신장비로 교체   하여 재 납품 하여야 한다.

4. 무상 및 유상 써비스 기간내에 원외반출 수리시에는 대체장비를 지원하여야 한다.

5. 장비납품 및 설치시 해당병원 장비별로 납품․설치하여야 하며, 정상운영 될 때까

지 초기교육을 실시하며 교육과 관련한 비용 및 장비설치에 따른 제반비용을 부담하여

야 한다.

6. 장비의 사용자의 변경이 생겼을 경우 Operator 교육을 실시하며 그 제반비용은 무상으

로 처리한다.

7. 무상 A/S기간 중 장비운영과 관련한 모든 프로그램의 Up-grade시 장비 제조사로부터

무상 제공하여야 한다.

8. 제품의 포장은 검수시에 개봉한 것만을 인정한다.

9. 납품은 병원이 요구하는 시간과 장소에 납품한다.

10. 장비의 검수는 장비설치 및 셋팅 완료 후에 실시한다.

11. 장비의 무상보증 기간은 장비 설치 및 검수 후부터 시작한다.

12. 사양서에 제시된 기능은 모두 사용이 가능하여야 하며 구성품 중 빠진 부분이 있   더라도 추가 납품한다.

13. 장비 설치시 병원의 설비의 변동이 없이 설치하는 것을 원칙으로 하되 부득이하게

공사 혹은 구조변경이 필요한 경우 병원에 사전에 통보하여야 하며 공사 시공시 병원

의 요청대로 시공하여야 한다. 또한 공사비용은 장비의 납품업체에서 부담한다.