사 양 요 구 서

부서명 : 산부인과

물품분류 번 호	품목번호	품명 (Descriptions)	단위	수량
		초음파기 (Ultrasound system)	set	1

1. 일반적인 요구사양

1. 산부인과용으로 최적화된 제품

2. 트랜듀서 연결포트가 4개 이상인 제품

3. 무상써비스기간이 3년 이상인 제품(트랜듀서포함)

4. Extension 모니터를 제공해 주는 제품

5. BW printer 및 color printer가 포함된 제품

6. 본원의 PACS와 연동이 가능한 제품

7. 모니터의 position 변동이 쉬운 제품

8. 220V, 60Hz의 전원을 사용하는 제품

9. 터치스크린으로 조작이 가능하며 사용자 편리하도록 설정이 가능한 제품

2. 사 양

1. Main Console

(1) System Operating/Display Modes

1) Digital Beamformer방식

2) System Processing channel

- Acquisition channel : 129,904 혹은 12,672

- Processing channel : 129,904 혹은 12,672

3) B-Mode, M-Mode,

4) Color Flow mode

5) PW Doppler with High PRF(PW)

6) Color Flow Imaging(CFI)

7) Power Doppler Imaging(PD)

8) Tissue Doppler Imaging(TD)

9) Simultaneous B/PW, B/CFM, B/PD, B/TD, B/M, B/3D

10) Real-time Triplex Mode: B/CFM/PW, B/PD/PW

11) Colorized B, M, and PW Doppler display

12) Imaging frequency range : 2 ~16 MHz 혹은  1.3 - 16.5 MHz

(2) B-Mode and M-Mode Imaging

1) Gray scale levels : 256 level

2) Persistence : 8 step

3) B Gray Scale Map : 12 maps 이상

4) B Edge enhancement : 5 step 이상

5) Line Filter : 3 step 이상

6) M Sweep Speed : 6 step 혹은 4 Type

7) User selectable dynamic range control : 200 dB (Wide DR) 혹은 180

dB

8) User selectable scan depth :  0cm - 30 cm 이상

9) B-mode Cine playback with storage

10) Colorized gray scale image capability  to highlight specific gray level

11) B Colorization : 5 chroma maps 이상

12) Write Zoom : 4 배이상

13) Read Zoom : 2.4배 이상

(3) CW- Mode  Doppler Imaging

1) Doppler Auto Trace

2) Base line : 0~8 Step

(4) Color Doppler Imaging 혹은 Color Flow Mode

1)  Frequency Range : Pen, Gen(depending on probe) 혹은  1-15Mhz

2) Color Maps : 8 Maps 이상

3) CDM Gain : 0 ~ 100 % 혹은 30dB range, 1dB steps

4) Line Density : 4 steps 혹은 10 steps

5) Line Filter : 4 steps 이상

6) CDM Spectrum Inversion

7) CDM Baseline Shift : 16 steps 이상

8) Balance : 16 steps 혹은 25-225,  step size 5

9) Color Display Mode :  V(Velocity),V-T(Velocity + Turbulence),

V-P(Velocity + Power), T(Turbulence), P-T(Power + Turbulence)

10) Realtime Triplex  Mode

- B + CDM/PW in any depth

(5) Power Doppler Imaging(PD)

1) Frequency Range : Pen, Gen(depending on probe) 혹은 1-15 Mhz

2) PD Maps : 8 Maps

3) Gain : 0~100% 혹은 30dB range, 1dB steps

4) Velocity Scale Range

- PRF : 1 ~ 23khz 혹은 100Hz to 11kHz

5) Wall Filter : 0 ~ 3 level 혹은 8 - 3000 Hz

6) Line Density : 4 steps 혹은 10 steps

7) Line Filter : 4 steps 혹은 8 steps

8) Maximum Steerable Angle : +/- 20 degree 혹은 +/- 25 degree or more

9) Balance : 1 ~ 16 steps 혹은 25-225,  step size 5

10) Realtime Triplex  Mode : B + CDM/PW in any depth 혹은 B +

CFM/PW in any depth

(6) Measurements and Calculations

1) General B-Mode and 3D

- Distance(Point to Point)

- 2D Trace(Trace to length)

2) Area/Circumference

- Ellipse/ Trace

3) Volume: following methods

- 1 Distance

- 1 Ellipse

- 1 Dist. + 1 Elipse

- 3 Distance

- Planimetric Volume(3D Only)

4) Hip Joint

5) Histogram

6) Comprehensive body pattern package

7) OB calculations and Fetal biometry with summary report

8) Vascular calculations with summary report and automatic trace

capability

9) Include fetal growth and trending with multi gestational growth and

gestational age

10) Cardio Measurements / calculations

11) Additional  Measurements / calculations

(7) Image /Volume Storage

1) On board image storage software

- Cine memory :  Image cine : 5400 frames, loop cine : 8192 lines 혹은

at least 3000 B-Mode images

- Dual/Quad Image CINE Display

(8) Image file

1) 3files, DICOM file(about 1MB), JPEG file(46KB)

2) On board MOD(Magneto optical Disk) for storing and retreival of

images supporting both 128MB,230MB, 520MB,640MB,1.3GB

3) Backup function to DVD/CD-R(W)

4) AVI files : conversion and expert to MO or DVD/CD-R(W)

5) Hard Disk Storage : 320 GB 이상

(9) Operator Interface

1) Active probe ports : 3 이상

2) Integrated HDD : 320GB 이상

3) Integrated DVD/CD-RW drive

4) Interactive back lighting keys

5) Context sensitive help utility

6) Touch Screen : 9inch 이상의 LED 혹은 LCD

7) Monitor : 19" LCD 이상

2. 이미지 최적화 기능

(1) QuickScan 혹은 Auto Optimization

(2) Tissue Harmonic 혹은 Coded Harmonics

(3) Trapezoidal 혹은 Virtual Convex

(4) Post Imaging Processing 혹은 Raw Data Processing

(5) SRF 혹은 SRI Iil

(6) DMR / DMR +  혹은 동등한 기능

(7) Dynamic Receive Focusing 혹은 FFC

(8) SCI 혹은 CRI( CrossXbeam )

(9) DPD ( Directional Doppler ) 혹은 HD-Flow

(10) Pulse Inversion 혹은 Inversion Mode

(11) 4D Live Imaging 혹은 Realtime 4D

(12) VOCAL 혹은 VOCALII

(13) Oblique View 혹은 Advaced VCI

(14) HDVI (HD Volume Imaging ) 혹은 동등한 기능

(15) Volume NT 혹은 동등한 기능

3. 트랜듀서

(1) Wide Band Convex Transducer

1) Frequency : 2.0 - 6.0 MHz 혹은 2.0 - 5.0 MHz

2) Application :  Abdomen, OB Gyn, Urology, Pediatrics.

(2) Wide Band Convex Volume Transducer

1)  Frequency : 4.0-8.0 MHz 혹은 2.0-8.0 MHz

2) Application : OB, GYN, Pediatric, Urology

(3) Wide Band Convex Transducer

(1) Frequency :  5.0 - 9.0 MHz 혹은  4.0 - 9.0 MHz

(2) Application : OB Gyn, Urology

(4) Phased array Transducers

(1) Frequency :  4.0 - 12.0 MHz 혹은  4.0 - 1.0 MHz

(2) Application : Cardiology, Pediatrics

4. DICOM3.0 Connectivity

3. 구성품

A) 혹은  B) 구성품

A) 구성품

1. 본체 ---------------------------------------------------1세트

(Harmonic Image, 3D Packge, 3DMXI, 이미지 저장장치, 키보드상 Touch 판넬,

FSI내장, SRF내장, SCI내장, Volume NT내장, Cadiac S/W, DICOM내장)

2. 19인치 LCD 모니터 --------------------------------------- 1ea

3. 2~6MHz 컨벡스 프로브(2D용) --------------------------------- 1ea

4. 4~9Mhz 바지날 프로브(2D용) VR5-9 -------------------------- 1ea

5. 4~8MHz 삼차원 컨백스 프로브(3D용) V4-8----------------------- 1ea

6. 4-12MHz 소아심장 프로브 (2D 용) P4-12 ----------------------- 1ea

7. HDVI USB ---------------------------------------------- 1ea

8. DMR PLUS USB ----------------------------------------- 1ea

9. CW ---------------------------------------------------- 1ea

10. DynamicMR plus ----------------------------------------- 1ea

11. HDVI ------------------------------------------------- 1ea

12. 3D MXI ------------------------------------------------ 1ea

13. DICOM 3.0 --------------------------------------------- 1ea

14. CW Functions ------------------------------------------- 1ea

15. Cardiac Measurement ------------------------------------- 1ea

16. ECG -------------------------------------------------- 1ea

17. 칼라 프린터 --------------------------------------------- 1ea

18. 흑백 프린터 --------------------------------------------- 1ea

19. 사용 설명서 --------------------------------------------- 1권

20. 서비스 매뉴얼 -------------------------------------------- 1권

21. Extension 모니터(19인치 이상) ------------------------------- 1ea

22. 기타 장비 구동에 필요한 악세사리 ---------------------------- 1set

B) 구성품

1. Main Console (CW Hardware) ------------------------------ 1set

2. Coded Excitation ----------------------------------------- 1ea

3. Auto Optimize ------------------------------------------- 1ea

4. Coded Harmonic Imaging ---------------------------------- 1ea

5. Virtual Convex ------------------------------------------ 1ea

6. True Access, Raw Data Processing(postprocessing) -------------- 1ea

7. FFC( Frequency Focus Composite) --------------------------- 1ea

8. CRI( CrossXbeam) -----------------------------------------1ea

9. HD-flow( High Definition Flow) ----------------------------- 1ea

10. Inversion Mode ----------------------------------------- 1ea

11. Real Time  4D ----------------------------------------- 1ea

12. Volume CALculation(VOCALII) ---------------------------- 1ea

13. VCI (Volume Contrast Imaging) ---------------------------- 1ea

14. DICOM3.0 Connectivity ----------------------------------- 1ea

15. Oranic Specific SRI III( Speckle Reduction Imaging) ------------ 1ea

16. TUI ( Tomographic Ultrasound Imaging) --------------------- 1ea

17. 19" size High resolution flat LCD monitor -------------------- 1ea

18. 4C-D Wide Band Convex Transducer ------------------------ 1ea

19. RAB4-8D Wide Band Convex Volume Transducer -------------- 1ea

20. IC5-9-D Wide Band Convex Transducer ----------------------1ea

21. PA6-8-D  Phased array Transducers ------------------------ 1ea

22. Color Printer ------------------------------------------- 1ea

23. B/W Printer ------------------------------------------- 1ea

24. Operating Manual --------------------------------------- 1ea

25. Extension 모니터(19인치 이상) ------------------------------ 1ea

26. 기타 장비 구동에 필요한 악세사리 ---------------------------- 1set

4. 기타사항

1. 상기 모든 내용을 모두 충족시키는 제품을 납품하여야 하며 납품 후에라도 상기

요구 사항에 맞지 않는 제품을 납품하였을 때에는 장비의 회수 및 결제 대금을 전액

변상하여야 하며 병원에 입은 손해에 대하여 보상하여야 한다.

2. 장비의 검수는 당 의료원의 의공담당자의 참석하에 실시 한 것만을 인정한다.

3. 장비설치 및 납품시 장비의 고장 및 불량이 발생하였을 경우 즉시 신장비로 교체   하여 재 납품 하여야 한다.

4. 무상 및 유상 써비스 기간내에 원외반출 수리시에는 대체장비를 지원하여야 한다.

5. 장비납품 및 설치시 해당병원 장비별로 납품․설치하여야 하며, 정상운영 될 때까

지 초기교육을 실시하며 교육과 관련한 비용 및 장비설치에 따른 제반비용을 부담하

여야 한다.

6. 장비의 사용자의 변경이 생겼을 경우 Operator 교육을 실시하며 그 제반비용은

무상으로 처리한다.

7. 무상 A/S기간 중 장비운영과 관련한 모든 프로그램의 Up-grade시 장비 제조사

로부터 무상 제공하여야 한다.

8. 제품의 포장은 검수시에 개봉한 것만을 인정한다.

9. 납품은 병원이 요구하는 시간과 장소에 납품한다.

10. 장비의 검수는 장비설치 및 셋팅 완료 후에 실시한다.

11. 장비의 무상보증 기간은 장비 설치 및 검수 후부터 시작한다.

12. 장비 설치시 병원의 설비의 변동이 없이 설치하는 것을 원칙으로 하되 부득이하

게 공사 혹은 구조변경이 필요한 경우 병원에 사전에 통보하여야 하며 공사 시공시

병원 의 요청대로 시공하여야 한다. 또한 공사비용은 장비의 납품업체에서 부담한

다.